## CALSHEET
Calibration Sheet
Test Stand No: |  | System 5000 |  | Mode: | Extending |  | Calibration Procedure No: |  |  |  | LMS-TD-5317-A
EX Voltage: | 2 | NIST Traceable Standard ID: |  |  |  | A022346 |  | V / FS: |  | mV Sensitivity: |  | -2.262
Xducer/Controller SN: |  | A024621 | 1874714 |  | Cal Range: | 0.150 in | Cal Sw:n/a |  | Sh Cal: | n/a
As Found |  |  |  | Corrected |  |  |  | As Left
STD | Machine | Error |  | STD | Machine | Error |  | STD | Machine | Error |  | Repeat
|  | % Reading |  |  |  | % Reading |  |  |  | % Reading |  | Error %
|  |  |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0
|  |  |  | 0.015 | 0.015 | 0 |  | 0.015 | 0.015 | 0 |  | 0
|  |  |  | 0.03 | 0.0299 | 0.33333 |  | 0.03 | 0.03 | 0 |  | 0.33333
|  |  |  | 0.045 | 0.0449 | 0.22222 |  | 0.045 | 0.0449 | 0.22222 |  | 0
|  |  |  | 0.06 | 0.0599 | 0.16667 |  | 0.06 | 0.06 | 0 |  | 0.16667
|  |  |  | 0.075 | 0.0749 | 0.13333 |  | 0.075 | 0.0749 | 0.13333 |  | 0
|  |  |  | 0.09 | 0.0899 | 0.11111 |  | 0.09 | 0.0899 | 0.11111 |  | 0
|  |  |  | 0.105 | 0.1049 | 0.095238 |  | 0.105 | 0.105 | 0 |  | 0.095238
|  |  |  | 0.12 | 0.12 | 0 |  | 0.12 | 0.12 | 0 |  | 0
|  |  |  | 0.135 | 0.135 | 0 |  | 0.135 | 0.135 | 0 |  | 0
|  |  |  | 0.15 | 0.15 | 0 |  | 0.15 | 0.1501 | -0.066667 |  | 0.066667
|  |  |  | 0.135 | 0.135 | 0 |  | 0.135 | 0.135 | 0 |  | 0
|  |  |  | 0.12 | 0.12 | 0 |  | 0.12 | 0.12 | 0 |  | 0
|  |  |  | 0.105 | 0.105 | 0 |  | 0.105 | 0.105 | 0 |  | 0
|  |  |  | 0.09 | 0.0899 | 0.11111 |  | 0.09 | 0.09 | 0 |  | 0.11111
|  |  |  | 0.075 | 0.075 | 0 |  | 0.075 | 0.0749 | 0.13333 |  | -0.13333
|  |  |  | 0.06 | 0.0599 | 0.16667 |  | 0.06 | 0.06 | 0 |  | 0.16667
|  |  |  | 0.045 | 0.0449 | 0.22222 |  | 0.045 | 0.0449 | 0.22222 |  | 0
|  |  |  | 0.03 | 0.03 | 0 |  | 0.03 | 0.03 | 0 |  | 0
|  |  |  | 0.015 | 0.015 | 0 |  | 0.015 | 0.015 | 0 |  | 0
|  |  |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0
Note: | Extensometer 256 | 2101 |  | Cal.By: | Will Johnston |  |  | Cal Date: | 2002-09-18 | Next Cal Due: |  | 2003-09-18